## Slide 1: Selection Of Water Generation Point

## Slide 2
1
Introduction
2
Types Of Water Sources in Industry
Page 3
Page 4
CONTENTS
Parameters for selection criateria
3
Page 10
About PMG Engineering
#
X
Required Testing For Water Source
4
Page 14

## Slide 3: Introduction
Water resources are sources of water that are potentially useful. All Living things required water to grow & reproduce.
They are respectively Surface water , Under river flow , Ground water , Frozen water , Desalination etc. Total quantity available of water at any time is important consideration.
Surface Water Water On the surface of the planet such as in a river, lake, Wetland or ocean.
Notes: 3

## Slide 4: Types Of Water Sources In Industry
Sources Of Water
  Desalination
  Desalination is an artificial process by which saline water (generally sea water) is converted to fresh water.
  The most common desalination processes are distillation and reverse osmosis.
  Under River Flow Water
  Total Volume of water transported downstream will often be a combination of the visible free water flow together with substantial contribution flowing through rocks and sediments that underlie the river and its floodplain called hyporheic zone.
  For many rivers in large valleys, this unseen component of flow may greatly exceed the visible flow. This is especially significant in karst areas where pot-holes and underground rivers are common.
Notes: 4

## Slide 5: Types Of Water Sources In Industry
Sources Of Water
  Ground Water
  Groundwater is fresh water located in the subsurface pore space of soil and rocks.
  It is also water that is flowing within aquifers below the water table.
  It is useful to make a distinction between groundwater that is closely associated with surface water and deep groundwater in an aquifer, sometimes called fossil water.
  Groundwater can be thought of in the same terms as surface water: inputs, outputs and storage.
  The critical difference is that due to its slow rate of turnover, groundwater storage is generally much larger (in volume) compared to inputs than it is for surface water.
Notes: 5

## Slide 6: Types Of Water Sources In Industry
Sources Of Water
  Frozen Water
  A water Several schemes have been proposed to make use of icebergs as a water source. However, to date this has only been done for research purposes. Glacier runoff is considered to be surface water.
Notes: 6

## Slide 7: Types Of Water Sources In Industry
Sources Of Water
  Borewells Or Tubewells
  Both are basically vertical drilled wells, bored into an underground aquifer in the earths surface, to extract water for various purposes.
  The difference in the two lies in the type of casing used, the depth of this casing and the type of soil where they are drilled.
  Casing to support the external surfaces of the borehole against collapse may be needed at certain depths, and usually is made up of PVC pipes.
  Electrical pumps are usually used to pump out the water from the borewells.
Notes: 7

## Slide 8: Types Of Water Sources In Industry
Sources Of Water
  Water Well
  A water well is an excavation or structure created in the ground by digging, driving, boring, or drilling to access groundwater in underground aquifers.
Canal Water
  Canals and navigations are human-made channels for water conveyance (supply), or to service water transport vehicles.
Notes: 8

## Slide 9: Parameters For Selection Criteria
  Parameters
TDS (Total Dissolve Solids)
  It refer to any minerals, salts, metals, cations or anions dissolved inwater.
  It comprise inorganic salts, principally calcium, magnesium, potassium, sodium, bicarbonates, chlorides, and sulfates and some small amounts of organic matter that are dissolved in water.
  Conductivity
  Conductivity is a measure of water's capability to pass electrical flow.
  This ability is directly related to the concentration of ions in the water 1.
  These conductive ions come from dissolved salts and inorganic materials such as alkalis, chlorides, sulfides and carbonate compounds.
Notes: 9

## Slide 10: Parameters For Selection Criteria
  Parameters
PH Value
  The pH of pure water is 7.
  In general, water with a pH lower than 7 is considered acidic, and with a pH greater than 7 is considered basic.
  The normal range for pH in surface water systems is 6.5 to 8.5, and the pH range for groundwater systems is between 6 to 8.5.
  BOD
  BOD stands on biological oxygen demand.
  It is the amount of dissolved oxygen needed by aerobic biological organisms to break down organic material present in a given water sample at certain temperature over a specific time-period.
Notes: 10

## Slide 11: Parameters For Selection Criteria
  Parameters
COD
  It stand on Chemical oxygen demand.
  It is an indicative measure of the amount of oxygen that can be consumed by reactions in a measured solution.
  It is commonly expressed in mass of oxygen consumed over volume of solution which in SI units is milligrams per litre(mg/L).
  TOC
  TOC stand on Total Oraganic Carbon .
  It is the amount of carbon found in an organic compound.
  It used as a non-specific indicator of water quality or cleanliness of pharmaceutical manufacturing equipment.
Notes: 11

## Slide 12: Parameters For Selection Criteria
  Parameters
Hardness
  Phosphate Calcium and magnesium dissolved in water are the two most common minerals that make water "hard.
  The hardness of water is referred to by three types of measurements: grains per gallon, milligrams per liter (mg/L), or parts per million (ppm).
Alkanity
  Alkalinity as the quantitative capacity of water to neutralize an acid that is, the measure of how much acid can be added to a liquid without causing a significant change in pH.
  pH more than 7 is known as alkline.
  Alkalinity is related to pH, because higher levels of alkalinity are useful in stabilizing the pH level.
Notes: 12

## Slide 13: Parameters For Selection Criteria
  Parameters
Turbidity
  It is a measure of the degree to which the water loses its transparency due to the presence of suspended particulates.
  More total suspended solids in the water, the murkier it seems and the higher the turbidity.
  Turbidity is considered as a good measure of the quality of water.
Nitrogen (Ammonia)
  Doesn’t react with water.
  Easily dissolves in water.
  Nitrogen (N2) solubility at 20oC and pressure = 1 bar is approximately 20 mg/L.
Notes: 13

## Slide 14: Required Testing For Water Source
  Testing for Total Dissolve Solids
Principle
  Are the residue left after the evaporation and drying in oven at 103-105°C.
  Include “Total suspended solids” (TSS) and “Total dissolved solids” (TDS).
  When same sample heated at 500°C, 50°C, organic matter is converted to CO2 which results in loss in weight (= volatile solids).
  Nitrogen, in the forms of nitrate, nitrite, or ammonium, is a nutrient needed for plant growth. Nitrogen gas does not react with water. It does dissolve in water. Nitrogen (N2) solubility at 20oC and pressure = 1 bar is approximately 20 mg/L.
Notes: 14

## Slide 15: Required Testing For Water Source
  Testing for Total Dissolve Solids
  Apparatus & Equipment
  Evaporatory dish (porcelain) – 100/200mL
  Drying oven – equipped with thermostatic control capable of maintaining the temperature within 2°C range.
  Desiccator – provided with desiccants
  Analytical balance – 200mg capacity of weighing to 0.1mg
  Filter holder – Gooch crucible adapter or membrane filters
  Suction flask – 500mL capacity
Notes: 15

## Slide 16: Required Testing For Water Source
  Testing for Total Dissolve Solids
  Procedure
  Filter the well-mixed sample under vacuum through membrane filter or Gooch Crucible.
  Transfer 100mL or more, depending upon the concentration of dissolved solids, in a weighed evaporating dish.
  Evaporate to dryness on steam bath.
  Dry the evaporated sample for at least 1 hour in an oven at 180±2°C.
  Cool in a desiccator and weigh.
  Repeat the drying until constant weight or weight loss less than 0.5mg is observed.
Notes: 16

## Slide 17: Required Testing For Water Source
  Testing for Total Dissolve Solids
Calculation:
  mg/L total filtrable residue at 180°C = (A –B) x 1000 / C
  Where,
    A = weight of dried residue + dish
    B = weight of dish
    C = mL of filtrate used
Notes: 17

## Slide 18: Required Testing For Water Source
  Testing for Conductivity
Principle
  This method is used to measure the conductance generated by various ions in the solution/water.
  Estimation of dissolved ionic contents of water equals to the product of specific conductance in µs/cm and empirical factor.
  Empirical factor vary from 0.55 to 0.90 depending on the soluble components of water and on the tempreture of measurements.
Apparatus and Equipment
  Conductivity meter
  Thermometer, capable of being read to the nearest 0.1°C and covering the range 10-50°C.
  Conductivity Cell –The cell choice will depend on expected range of conductivity and the resistance range of instrument.
Notes: 18

## Slide 19: Required Testing For Water Source
  Testing for pH Value
Principle
  The pH is determined by measurement of the electromotive force of a cell comprising of an indicator electrode immersed in the test solution and a reference electrode.
  Contact is achieved by means of a liquid junction, which forms a part of the reference electrode. The emf of this cell is measured with pH meter.
  Since the pH is defined operationally on a potentiometric scale, the measuring instrument is also calibrated potentiometrically with an indicating (glass) electrode and a reference electrode using standard buffers having assigned pH value.
Notes: 19

## Slide 20: Required Testing For Water Source
  Testing for pH Value
The pH (B) = -log10 [H+]
where pH (B) = assigned pH of standard buffer.
The operational pH scale is used to measure sample pH and is defined as:
pHs = pH (B) + F (Es – EB) / 2.303 RT
Where,
  pHs = potentiometrically measured sample Ph
  F = Faraday 9.649 x 104 coutomb/mole
  Es = Sample emf V
  EB = Buffer emf V
  R = Gas constant 1.987 cal deg-1 mole-1
  T = absolute temperature, °K
Notes: 20

## Slide 21: Required Testing For Water Source
  Testing for Conductivity
Apparatus and Equipement
  pH meter consisting of potentiometer, glass electrode, referance electrode and tempreture compensating device.
  Referance electrode consisting of half cell that provides a standard electrical potential.
  Sensor glass electrode
  Beakers preferably use polyethylene or TFE beakers.
  Used magnetic TFE coated stirring bar.
Procedure
  Before use, remove electrodes from storage solutions and rinse with distilled water.Dry electrodes by gently blotting with a soft tissue paper, standardise instrument with electrodes immersed in a buffer solution within 2 pH units of sample pH.
  Remove electrodes from buffer, rinse thoroughly with distilled water and blot dry.
Notes: 21

## Slide 22: Required Testing For Water Source
  Testing for Conductivity
Procedure
  Immerse in a second buffer below pH 10, approximately 3 pH units different from the first, the reading should be within 0.1 unit for the pH of second buffer.
  For samples analysis, establish equilibrium between electrodes and sample by stirring sample to ensure homogeneity and measure pH.
  For buffered samples (or those with high ionic strength), condition the electrodes after cleaning by dipping them into the same sample, and read pH.
  With poorly buffered solutions (dilute), equilibrate electrodes by immersing in three or four successive portions of samples. Take a fresh sample and record the pH.
Calculation
  pH value is obtained directly from Instrument.
Notes: 22

## Slide 23: Required Testing For Water Source
  Testing for BOD
Principle
  This test measures the oxygen utilised for the biochemical degradation of organic material (carbonaceous demand) and oxidation of inorganic material such as sulphides and ferrous ions during a specified incubation period.
  It also measures the oxygen used to oxidize reduced forms of nitrogen (nitrogenous demand) unless their oxidation is prevented by an inhibitor.
  The methodology of BOD test is to compute a difference between initial and final DO of the samples incubated.
  Minimum 1.5 L of sample is required for the test.
Notes: 23

## Slide 24: Required Testing For Water Source
  Testing for BOD
  Apparatus & Equipment
  BOD bottles 300ml capacity with a water seal.
  Incubator or water-bath to be controlled at 20ºC or at any desired temperature 1ºC.
  Procedure
  The source of dilution water may be distilled water, tap or receiving-stream water free of biodegradable organics and bioinhibitory substances such as chlorine or heavy metals.
  Aerate the required volume of dilution water in a suitable bottle by bubbling clean-filtered compressed air for sufficient time to attain DO saturation at room temperature or at 20ºC/27ºC. Before use stabilise the water at 20ºC/27ºC.
Notes: 24

## Slide 25: Required Testing For Water Source
  Testing for BOD
Add 1mL each of phosphate buffer, magnesium sulphate, calcium chloride and ferric chloride solutions in that order for each Litre of dilution water. Mix well.
Quality of dilution water may be checked by incubating a BOD bottle full of dilution water for 5 days at 20ºC for 3 days at 27ºC. DO uptake of dilution water should not be more than 0.2mg/L and preferable not more than 0.1mg/L.
Calculation:
  When Dilution water is not seeded
  BOD as O2 mg/LI = (D1-D2)X100
          %Dilution
When Dilution is seeded
          BOD as O2 mg/Ll= (D1-D2)-(B1-B2)X100
          %Dilution
Notes: 25

## Slide 26: Required Testing For Water Source
  Testing for BOD
  Calculation:
  When Dilution water is not seeded
  BOD as O2 mg/Ll = (D1-D2)X100
          %Dilution
          When dilution is seeded
  BOD as O2 mg/Ll= (D1-D2)-(B1-B2)X100
          %Dilution
          Where,
            D1 = DO of sample immediately after preparation, mg/L
            D2 = DO of sample after incubation period, mg/L
            B1 = DO of blank (seeded dilution water) before incubation, mg/L
            B2 = DO of blank (seeded dilution water) after incubation, mg/L
Notes: 26

## Slide 27: Required Testing For Water Source
  Testing for COD
Principle
  The dichromate reflux method is preferred over procedures using other oxidants because of its superior oxidizing ability, applicability to a wide variety of samples and ease of manipulation. Oxidation of most organic compounds is up to 95-100% of the theoretical value.
  The organic matter gets oxidised completely by potassium dichromate with silver sulphate as catalyst in the presence of concentrated H2SO4 to produce CO2 and H2O. The excess K2Cr2O7 remaining after the reaction is titrated with ferrous ammonium sulphate [Fe(NH4)2(SO4)2].
Notes: 27

## Slide 28: Required Testing For Water Source
  Testing For BOD
  Apparatus & Equipment
  250 or 500mL Erlenmeyer flask with standard (24/40) tapered glass joints.
  Friedrich reflux condenser(12 Inch) with standard (24/40) tapered glass joints.
  Electric hot plate or six-unit heating shelf.
  Volumetric pipettes (10, 25, and 50mL capacity).
  Burette, 50mL with 0.1mL accuracy.
  Burette stand and clamp.
  Analytical balance, accuracy 0.001g.
  Spatula And Volumetric flasks (1000mL capacity)
Notes: 28

## Slide 29: Required Testing For Water Source
  Testing for COD
  Procedure
  All samples high in solids should be blended for 2 minutes at high speed and stirred when an aliquot is taken for analysis.
  Select the appropriate volume of sample based on expected COD range, e.g. for COD range of 50-500 mg/L take 25-50mL of sample.
  Sample Volume less than 25mL should not be pipetted directly, but serially diluted and then a portion of the diluted sample taken.
  Dilution factor should be incorporated in calculation.
Notes: 29

## Slide 30: Required Testing For Water Source
  Testing for COD
  Calculation
  COD as mg/L = (a –b) x N x 8000 / mL sample
  Where;
    a = mL FAS used for blank
    b = mL FAS used for sample
Notes: 30

## Slide 31: Required Testing For Water Source
  Testing for TOC
Principle
  Instrumental TOC method utilises high temperature, catalysts and oxygen or lower temperature (<100°C) with irradiation, chemical oxidants, or combinations of these oxidants to convert organic carbon to carbon dioxide (CO2). The CO2 may be measured directly by a non-dispersive infrared analyser or CO2 may be titrated chemically.
  Eliminating or compensating for IC interferences requires multiple determinations to measure TOC. IC interference can be eliminated by acidifying samples to pH 2 or less to convert IC species to CO2.
  purging the sample with a purified gas removes the CO2 by volatilization. Alternatively, IC interference may be compensated for by separately measuring total carbon (TC) and inorganic carbon (IC).
  The difference between TC and IC is TOC.
Notes: 31

## Slide 32: Required Testing For Water Source
  Testing for TOC
  Apparatus & Equipment
  Total organic carbon analyzer.
  Syringes: 0 to 50μL, 0 to 200μL, 0 to 500μL, and 0 to 1mL.
  Sample blender or homogenizer.
  Magnetic stirrer and TFE-coated stirring bars.
  Filtering apparatus and 0.45μm-pore diameter filters
Notes: 32

## Slide 33: Required Testing For Water Source
  Testing for TOC
  Procedure
  Take standard solution in the suitable range.
  Inject standard for calibration curve.
  Run a blank.
  Vary the sample size from 200 to 100 µL.
  Sample free from inorganic carbon, reduce the pH upto 2 and pass Co2.
  Take a sample from these by a capillary syringe and inject into the analyser.
  Measure the response and repeat with three such estimation to chain an average.
Notes: 33

## Slide 34: Required Testing For Water Source
  Testing for TOC
  Calculation:
  The instrument provides the value for TC and IC for sample.
  Calculate TOC By substraction.
  Express the result mg/L.
Notes: 34

## Slide 35: Required Testing For Water Source
  Testing for Hardness
Principle
  Hardness is determined by the EDTA method in alkaline condition.
  When EDTA is added as a titrant, calcium and magnesium divalent ions get complexed resulting in sharp changes from wine red to blue which indicates end point of the titration.
  Magnesium ion must be present to yeild satisfactory point of the titration.
  Small amount of complexomatically neutral magnesium salt of EDTA is added to the buffer.
  The sharpness of end point increases with increasing pH.
Notes: 35

## Slide 36: Required Testing For Water Source
  Testing for Hardness
Apparatus and Equipment
  Conical flask 100 mL
  Burette
  Pipette
  Spatula
  Procedure
  Take 25 or 50mL well mixed sample in porcelain dish or conical flask.
  Add 1-2mL buffer solution followed by 1mL inhibitor.
  Add a pinch of Eriochrome black T and titrate with standard EDTA (0.01M) till wine red colour changes to blue, note down the volume of EDTA required conical flask.
Notes: 36

## Slide 37: Required Testing For Water Source
  Testing for Hardness
Run a reagent blank.
Calculate volume of EDTA required by sample, C = (A-B).
For natural waters of low hardness, take a larger sample volume, i.e., 100-1000mL for titration and add proportionately larger amounts of buffer, inhibitor and indicator .
Add standard EDTA titrant slowly from a micro burette and run a blank using redistilled, deionised water of the same volume as sample.
Notes: 37

## Slide 38: Required Testing For Water Source
  Testing for Hardness
  Calculation:
  Total hardness as CaCO3 mg/L =CXDX1000/mL Sample
  Where,
    C= Volume of EDTA required by sample .
    D= mg CaCO3 equivalent to 1 mL EDTA titrant.
Notes: 38

## Slide 39: Required Testing For Water Source
  Testing for Alkanity
Principle
  Hardnes Alkalinity of sample can be estimated by titrating with standard sulphuric acid (0.02N) at room temperature using phenolphthalein and methyl orange indicator.
  Titration to decolourisation of phenolphthalein indicator will indicate complete neutralization of OH- and ½ of CO3-, while sharp change from yellow to orange of methyl orange indicator will indicate total alkalinity.
Apparatus and Equipment
  Beakers – size and form will depend on electrode and the size of the sample to be used for determination of alkanity.
  Pipittes ( Volumetric)
  Flasks (Volumetric ) 1000 mL, 200mL,100mL.
Notes: 39

## Slide 40: Required Testing For Water Source
  Testing for Alkanity
Procedure
  Take 25 or 50mL sample in a conical flask and add 2-3 drops of phenolphthalein indicator.
  If pink colour develops titrate with 0.02N H2SO4 till disappears or pH is 8.3. Note the volume of H2SO4 required.
  Add 2-3 drops of methyl orange to the same flask and continue titration till yellow colour changes to orange. Note the volumes of H2SO4 required.
  In case pink colour does not appear after addition of phenolphthalein continue as above.
  Alternatively, perform potentiometric titration to preselected pH using appropriate volume of sample and titration assembly. Titrate to the end point pH without recording intermediate pH.
Notes: 40

## Slide 41: Required Testing For Water Source
  Testing For Alkanity
Calculation
  Calculate total (T), phenolphthalein (P) alkalinity
  P-alkalinity, as mg CaCO3/L = A x 1000/mL sample
  T-alkalinity, as mg CaCO3/L = B x 1000/mL sample
  In case H2SO4 is not 0.02 N apply the following formula
    Alkalinity, as mg CaCO3/L = A/B x N x 50000 / mL of sample
    Where,
    A = mL of H2SO4 required to bring the pH to 8.3
    B = mL of H2SO4 required to bring the pH to 4.5
    N = normality of H2SO4
Notes: 41

## Slide 42: Required Testing For Water Source
  Testing for Turbidity
Principle
  Measured By its effect on the scattering light, which is termed as Nephelometry.
  Turbidimeter can be used for sample with moderate turbidity and nepelometer for sample with low turbidity.
  Higher the intensity of scattered lights higher the turbidity.
  Turbidity is an expression of the optical property that causes light to be scattered and absorbed rather than transmitted in straight lines through the sample.
  Standard method for the determination of turbidity has been based on the Jackson candle turbidity meter.
  The lowest turbidity value that can be measured directly on above instrument is 25 units.
  Indirect method is necessary to estimate the turbidity in the range of 0-5 units.
Notes: 42